Universal Console

- Should have in built user friendly interactive menu and illustrative help guide

- The various parameters should be able to adjust either from touch screen panel or from the multifunction foot switch

- Should be able to connect multiple hand pieces at a time like  Debrider hand pieces (Up to 7500 RPM in Oscillating mode & 30000 RPM in Forward mode) and drills (Up to 80000 RPM),

- Console should recognize the various hand pieces and automatically adjust the settings accordingly

- Should have inbuilt pumps each for Irrigation(5Cc / Min to 100Cc / Min)  and Cooling

- Should have multifunction ergonomically designed foot control with light emission for easy identification

- Should be able to control Speed / Mode, Forward / Reverse toggle, Active hand piece change from the Foot control itself

- Should have option for remote control Irrigation to operate from sterile area

- Should have in built scope Lens cleaning system.

- Should have provision to connect facial nerve monitor with the Otologic drill.

- Should have the option to work without foot paddle

Microdebrider Handpiece :

- Able to work up to the speed of 30000 RPM in forward rotation and 5000RPM in oscillation mode.

- Have fingertip control to rotate only the tip of the blade up to 360 deg.

- Have straight suction path to reduce clogging and allow efficient tissue removal.

- Have integrated blade locking system to lock the blade tip rotation.

- Have integrated side grooves and cable clips to provide better tubing management.

- Have Titanium body to avoid rusting.

- Should come with aZero, 12, 40,60,90 and 120 degree blades

Microdrill Specs

- Drill handpiece should be ergonomically designed and easy to hold while operating.

- Drill must be useful for lateral skull base as well as transnasal skull base procedures.

- The system should have electric motor which has maximum stall torque more than 42 mNM.

- Motor should run upto maximum speed of 75000 RPM .

- Weight of the drill should not be more than 90 grams and length should be less than 11.0 cm with diameter not exceeding 1.70 cm.

- Handpiece should have quick release and lock system for tools and attachment.

- The handpiece should be compatible with trans nasal skull base burrs which should be directly attachable to the motor.

- Must have option to connect irrigation tubing which can be programmed to provide intermittent or continuous irrigation.